APPROVED by

CEO

TestGene LLC

______________A. Toropovsky

"26" August 2021

INSTRUCTION FOR USE

Reagent kit for simultaneous qualitative detection of NA viruses of hepatitis B, C, D by multiplex RT-PCR

“HEPA-BCD-test”

Technical specifications 21.20.23-021-97638376-2020

Version 2 from 08/26/2021

Table of contents

Introduction	4

- Purpose	6

- Principle of the method	6

- Composition of the reagent kit	9

- Characteristics of the reagent kit	eleven

- List of risks associated with the use of the kit reagents	19

- Precautions when working with the kit	20

- Equipment and materials required when working with

set of reagents	22

- Analyzed samples	23

- Preparing test kit components	28

- Carrying out analysis	28

- Registration and interpretation of results	31

- Conditions of storage, transportation and operation

set of reagents	33

- Disposal	35

- Warranty obligations, contacts	36

List of abbreviations

The following abbreviations and designations apply in these instructions:

| PCR | polymerase chain reaction |
| --- | --- |
| RT | reverse transcription |
| DNA | deoxyrubonucleic acid |
| cDNA | complementary DNA |
| RNA | ribonucleic acid |
| HBV | hepatitis B virus |
| HCV | hepatitis C virus |
| HDV | hepatitis D virus |
| ICS | internal control sample |
| NCS | negative control sample |
| PCS | positive control sample |
| CSS | control sensitivity sample |
| SBS | specificity control sample |

Introduction

Infection with hepatitis B and C viruses is the main cause of acute and chronic liver disease. The hepatitis D virus, despite its lower prevalence, when co-infected or superinfected with hepatitis B, significantly complicates the course of the disease. Testing and timely diagnosis of hepatitis determine the indications for adequate therapy and allow monitoring of antiviral therapy.

Target analyte:specific sections of the genomic DNA of the hepatitis B virus (hepatitis B, HBV), RNA of the hepatitis C viruses (hepatitis C, HCV) and D (hepatitis delta, HDV).

Scientific validity of the target analytelies in its specificity (the uniqueness of the DNA or RNA sequence) in relation to the genomes of the hepatitis B, C and D viruses.

Hepatitis B virus is a member of the hepadnaviridae family. The genome is represented by a circular DNA molecule.

Hepatitis C virus is a member of the flavivirus family (Flaviviridae). The genome is represented by a single-stranded linear RNA molecule. An important feature of the genome is high heterogeneity. To date, there are 8 types.

Hepatitis D virus is a member of the genus Deltavirus (family undefined). The genome is represented by a circular single-stranded RNA molecule about 1700 nucleotides long. The presence of hepatitis B virus is required for viral replication.

Hepatitis B virus (HBV) and Hepatitis C virus (HCV) infections remain a global health problem. In the Russian Federation, the incidence of hepatitis B virus infection ranges from 1.85 to 128.69 per 100 thousand population in different regions of the Russian Federation. In 2018, 61.9 thousand cases of infection with the hepatitis C virus were detected in Russia. This clearly illustrates the relevance of diagnosing hepatitis viruses. The hepatitis D virus, despite its lower prevalence, when co-infected or superinfected with hepatitis B, significantly complicates the course of the disease.

Scope of application of the reagent kit:clinical laboratory diagnosis of infectious diseases.

Indications and contraindications for use

Indications for use: if infection with viral hepatitis B, C and D is suspected.

Contraindications for use: when used by specially trained personnel and taking into account the intended use, no contraindications have been identified.

Population and demographic aspects of the use of a medical device:No population or demographic aspects of the use of the “HEPA-BCD-test” reagent kit have been identified.

Sterility:the product is not sterile.

  - Intended use

Intended use:The “HEPA-BCD-test” reagent kit is intended or qualitative detection and differentiation of hepatitis B, C and D viruses using the one-step reverse transcription –multiplex allele-specific polymerase chain reaction with real-time hybridization-fluorescence detection (RT-PCR) in a sample of nucleic acids isolated from human K2-EDTA blood plasma, in patients with suspected infection with viral hepatitis B, C and D.

Functional purpose: the results obtained can be used to diagnose diseases caused by hepatitis B, C and D viruses.

Potential consumers of a medical product

The set is intended for professional use in medical institutions and clinical diagnostic laboratories. Professional level of potential users – clinical laboratory diagnostics doctor, medical laboratory technician, laboratory technologist.

  - Principle of the method

Method

One-step reverse transcription – multiplex allele-specific polymerase chain reaction with real-time hybridization-fluorescence detection (RT-PCR).

Type of sample analyzed

The material for RT-PCR is nucleic acid samples isolated from human K2-EDTA blood plasma.

Determination principle

Qualitative detection of specific sections of the genomic DNA of the hepatitis B virus and RNA of the hepatitis C and D viruses by the method of one-step reverse transcription - multiplex allele-specific polymerase chain reaction with hybridization-fluorescent detection in real time (RT-PCR) in a nucleic acid sample (hereinafter – NA), isolated from clinical material, includes three stages:

- RT-PCR preparation;

- Reverse transcription of RNA and PCR amplification of DNA with hybridization-fluorescence detection in real time;

- Interpretation of results.

With NA samples, a one-step reverse transcription reaction is carried out - a multiplex allele-specific polymerase chain reaction of specific regions of the genome of hepatitis viruses using primers in the reaction buffer. If hepatitis C and D viruses are present in the sample, their RNA is reversed during reverse transcription to form cDNA; If the hepatitis B virus is present in the sample, its DNA goes through the stage of reverse transcription without changes. Upon completion of reverse transcription, a “hot start” is carried out, as a result of which reversease is inactivated and polymerase is activated, and the stage of PCR amplification of DNA and cDNA is launched.

The PCR buffer contains all the basic reagents, including warm-start reversease, hot-start thermostable DNA polymerase, deoxynucleotide triphosphates and an optimized buffer.

Mixtures of oligonucleotides contain fluorescently labeled oligonucleotide probes, which hybridize with the complementary region of the target DNA being amplified and are hydrolyzed (destroyed) by Taq polymerase, as a result of which the dye and quencher are separated, and the fluorescence intensity increases in the corresponding range of the optical spectrum. This allows the accumulation of a specific amplification product to be detected by measuring the intensity of the fluorescent signal in real time.

The kit contains reagents for multiplex determination of highly specific regions (targets) of the genomic DNA of the hepatitis B virus and RNA of the hepatitis C and D viruses, as well as an internal control sample (ICS) (Table 1).

Table 1 - Analyzed targets

| Channel corresponding to fluorophore |  |  |  |
| --- | --- | --- | --- |
| FAM/Green | HEX/Yellow | ROX/Orange | Cy5/Red |
| HBV DNA | ICS | HСV RNA | HDV RNA |

ICS allows you to assess the quality and efficiency of NA isolation and the possible presence of reverse transcription and amplification inhibitors in the sample, the presence of which can lead to false negative results.

Limitations of the method

A possible reason for obtaining a false positive result is contamination at the stage of NA isolation or multiplex RT-PCR/PCR reaction. A false positive result can be detected by using a negative control sample.

Violation of the integrity of the packaging during transportation.

Using an expired kit or violating the storage conditions of the kit.

Violation of storage conditions during transportation of samples.

The conclusion about the clinical diagnosis cannot be based only on the results of the study using this MI. For diagnostic purposes, the results should be used in combination with other data: symptoms, general clinical picture, results of other test systems (for example, serological tests designed to detect antibodies or antigens using enzyme-linked immunosorbent assay or chemiluminescence immunoassay), therapy used.

If the viral load is very low (less than 250 copies/ml), which may be due to the antiviral therapy used or the course of the disease, false negative results may be obtained. In these cases, it is recommended to use more clinical

material for NA isolation to lower the sensitivity threshold of the test system.

The RT-PCR time ranges from 120 to 145 minutes (excluding sample preparation), depending on the cycler used.

  - Composition of the reagent kit

Set of reagents "HEPA-BCD test» is produced in one configuration form – «HEPA-BCD test"

Number of analyzed samples

The set of reagents is designed for 96 reactions, which corresponds to the determination of 94 test samples, negative and positive samples with a single run of a 96-well amplifier or 32 single setups of test samples with negative and positive control samples in each setup.

Table 2 – Composition of the “HEPA-BCD-test” reagent kit

| № | Reagent name | Description | Quantity, volume |
| --- | --- | --- | --- |
| 1. | PCR-buffer | Transparent colorless liquid | 1 test tube, 480 µl |
| 2. | Oligonucleotide mixture | Transparent colorless liquid, may have a tinge of lilac colour | 1 test tube, 1,440 µl |
| 3. | PCS | Transparent colorless liquid | 1 test tube, 160 µl |
| 4. | NCS | Transparent colorless liquid | 2 test tubes, 1,600 µl each |
| 5. | ICS | Transparent colorless liquid | 1 test tube, 950 µl |

Note: Operating documentation (instructions for use and quality certificate) are not included in the product, but are included in the product delivery package. A set of reagents, to ensure compliance with transportation conditions, is placed in a reusable polyurethane foam thermal container for temporary storage and transportation with prepared cold elements. The thermal container is placed in a cardboard box and placed there

instructions for use and quality certificate for each supplied batch of product.

Set contents

The PCR-buffer contains all the basic reagents, including warm-start reversease, hot-start thermostable DNA polymerase, deoxynucleotide triphosphates and an optimized buffer.

The oligonucleotide mixture is ready to use and contains primers and probes designed to identify specific targets - see Table 1. The oligonucleotide mixture is in a 10% aqueous solution of TE (1 mM Tris, 0.1 mM EDTA), free of nucleases.

PCS ready for use and is a mixture of plasmid DNA with synthetic inserts of amplified fragments of the genomic DNA of the hepatitis B virus and cDNA of the hepatitis C and D viruses, as well as a fragment of the bacteriophage genome. PCS is in 10% TE buffer (1 mM Tris, 0.1 mM EDTA).

NCS ready to use and is DNase and RNase free deionized water.

ICS is a preparation of reinforced RNA in TE buffer (10 mM Tris, 1 mM EDTA).

The kit does not contain drugs for medical use, substances of human or animal origin.

  - Characteristics of the reagent kit

  - Technical and functional characteristics

Table 3 – “HEPA-BCD-test” reagent kit

| Indicator name | Characteristics and standards | Item TU |
| --- | --- | --- |
| 1. Specifications |  | 1 |
| 1. Appearance |  |  |
| PCR-buffer | Transparent colorless liquid | Section 7 paragraph 7.6 |
| Oligonucleotide mixture | Transparent colorless liquid, may have a lilac tint | Section 7 paragraph 7.6 |
| PCS | Transparent colorless liquid | Section 7 paragraph 7.6 |
| NCS | Transparent colorless liquid | Section 7 paragraph 7.6 |
| ICS | Transparent colorless liquid | Section 7 paragraph 7.6 |
| 1.2. Completeness | In accordance with clause 1.4 of TS 21.20.23-021-97638376-2020 | Section 7 paragraph 7.12 |
| 1.3. Marking | In accordance with clause 4 of TS 21.20.23-021-97638376-2020 | Section 7 paragraph 7.12 |
| 1.4. Package | In accordance with clause 5 of TS 21.20.23-021-97638376-2020 | Section 7 paragraph 7.12 |
| 2. Functional characteristics |  |  |
| 2.1Positive result with PCS | Registration of fluorescence signal growth in tubes with PCS via FAM channels Сt≤25, HEX Сt≤25, ROX Сt ≤25 and Cy5 Сt ≤25. | Section 7 clause 7.8.2 |
| 2.2 Negative result with NCS | In test tubes with NCS, Ct is not indicated for the FAM, HEX, ROX and Cy5 channels (that is, there is no fluorescence accumulation graph) | Section 7 clause 7.8.2 |
| 2.3 Reaction in test tubes with SBS | In test tubes with SBS Ct through FAM channels, ROX and Cy5 are not specified (that is, the schedule there is no accumulation of fluorescence), and through the HEX channel Ct≤32. | Section 7 clause 7.8.2 |
| 2.4 Reaction in test tubes with CSS | In test tubes with CSS via FAM, ROX channels and Cy5 in all repeats (at least 4) Ct≤35, and in the HEX channel Ct≤32. | Section 7 clause 7.8.2 |

Note: When carrying out control PCR, the following are used as CSS and SBS:

- The control sensitivity sample-B (CSS-B) is a mixture of plasmids with synthetic inserts fragment of genomic DNA of the hepatitis B virus with a concentration of 250 copies/ml in TE buffer (10 mM Tris, 1 mM EDTA).

- The control sensitivity sample-C (CSS-C) is a mixture of plasmids with synthetic inserts of a fragment of the genomic cDNA of the hepatitis C virus (HCV_1pos) with a concentration of 250 copies/ml in TE buffer (10 mM Tris, 1 mM EDTA).

- The control sensitivity sample-D (CSS-D) is a mixture of plasmids with synthetic inserts of a fragment of the genomic cDNA of the hepatitis D virus (HDV_pos) with a concentration of 250 copies/ml in TE buffer (10 mM Tris, 1 mM EDTA).

- Specificity control sample (SCS), which is a solution of human genomic DNA isolated from the Jurkat cell line with a concentration of 2,000 copies per 5 μl (400,000 copies/ml) in TE buffer (10 mM Tris, 1 mM EDTA).

Samples CSS-B, CSS-C, CSS-D and KOS in a volume of 100 μl undergo the stage of NC isolation with the addition of 10 μl of ICS according to the instructions.

  - Analytical performance characteristics

    - Analytical Specificity

Specific to hepatitis B virus DNA and hepatitis C and D virus RNA.

The possibility of detecting RNA of the hepatitis D virus and various genotypes of hepatitis B (A, B, C, D, E, F, G) and C (from 1a to 6) using standard NIBSC samples has been confirmed:

- International standard 1st WHO International Reference Panel for Hepatitis B Virus Genotypes for Nucleic Acid Amplification Techniques -Based Assays PEI code 5086/08, consisting of 15 samples of lyophilized HBV-positive plasma and covering the most common HBV genotypes: Samples 1–3 (genotype A) , Samples 4–6 (genotype B), Samples 7–9 (genotype C), samples 10–12 (genotype D), sample 13 (genotype E), sample 14 (genotype F) and sample 15 (genotype G).

- 4th HCV RNA Genotype Panel for Nucleic Acid Amplification Techniques. NIBSC code: 14/290, consisting of seven bottles representing six main genotypes: bottle 14/276 – HCV genotype 1a, 14/278 – HCV genotype 1b, 14/280 – HCV genotype 2i, 14/282 – HCV genotype 3a, 14/284 – HCV4r genotype, 14/286 – HCV genotype 5a, 14/288 – 6l genotype.

- International standard 1st World Health Organization International Standard for Hepatitis D Virus RNA for Nucleic Acid Amplification Techniques (NAT)-Based Assays PEI code 7657/12, with a concentration of 575,000 IU/ml.

Based on the results of the study with standard NIBSC samples, the specificity of the “HEPA-BCD-test” reagent kit was confirmed in relation to the DNA of the hepatitis B virus (genotypes A, B, C, D, E, F, G) and the RNA of hepatitis C viruses (genotypes from 1a up to 6) and D.

    - Analytical sensitivity (detection limit)

In accordance with GOST R 51352-2013 and taking into account international recommendations CLSI EP-17A2, the limit of detection (LOD) was determined by dilution analysis of WHO International Standards:

- WHO International Standard 4th WHO International Standard for HBV DNA for NAT NIBSC code: 10/266, with a concentration of 955,000 IU/ml (~ 5.98 log10 IU/ml)

- International Standard WHO International Standard 6th WHO International Standard for hepatitis C virus RNA for nucleic acid amplification techniques NIBSC code: 18/184, with a concentration of 257,000 IU/tube (~ 5.41 log10 IU/tube)

- International standard 1st World Health Organization International Standard for Hepatitis D Virus RNA for Nucleic Acid Amplification Techniques (NAT)-Based Assays PEI code 7657/12, with a concentration of 575,000 IU/ml

in HBV/HCV/HDV-negative human K2-EDTA plasma within the expected detection limit for 100 µl samples.

Each of the 7 dilutions was analyzed using a kit

"HEPA-BCD test»on 3 different days in 30 replicates to calculate the percentage of positive results. The results were determined in accordance with international recommendations CLSI EP-17A2 by probit analysis.

Analytical sensitivity: the limit of detection (LOD) for 100 μl K2-EDTA human plasma samples was:

| Analyst | Cycler used | Concentration, IU/ml (LOD) s confidence level 95% | Confidence interval 95% |
| --- | --- | --- | --- |
| HBV DNA | DTprime | 47.8 | 95%CI: 46.37-49.23 |
|  | CFX 96 | 48.4 | 95%CI: 46.97-49.83 |
|  | Rotor-Gene Q | 48.1 | 95%CI: 46.67-49.53 |
|  | Quant Studio 5 | 48.8 | 95%CI: 47.37 – 50.23 |
| HCV RNA | DTprime | 96.7 | 95%CI: 95.2-98.1 |
|  | CFX 96 | 97.0 | 95%CI: 95.5-98.4 |
|  | Rotor-Gene Q | 105.9 | 95%CI: 104.4-107.3 |
|  | Quant Studio 5 | 97.1 | 95%CI: 95.6-98.5 |
| HDV RNA | DTprime | 39.2 | 95%CI: 37.7-40.63 |
|  | CFX 96 | 38.8 | 95%CI: 37.3-40.2 |
|  | Rotor-Gene Q | 39.4 | 95%CI: 37.9-40.8 |
|  | Quant Studio 5 | 36.5 | 95%CI: 35.1-37.9 |

To verify the detection limit when testing various genotypes of the hepatitis B virus (A, B, C, D, E, F, G) and hepatitis C virus (from 1a to 6), samples from the NIBSC standard panel were used:

- International standard 1st WHO International Reference Panel for Hepatitis B Virus Genotypes for Nucleic Acid Amplification Techniques-Based Assays PEI code 5086/08, consisting of 15 samples of lyophilized HBV-positive plasma and covers the most common HBV genotypes: Samples 1–3 (genotype A) , Samples 4–6 (genotype B), Samples 7–9 (genotype C), samples 10–12 (genotype D), sample 13 (genotype E), sample 14 (genotype F) and sample 15 (genotype G).

- 4th HCV RNA Genotype Panel for Nucleic Acid Amplification Techniques. NIBSC code: 14/290, consisting of seven bottles representing six main genotypes: bottle 14/276 – HCV genotype 1a, 14/278 – HCV genotype 1b, 14/280 – HCV genotype 2i,

14/282 – HCV genotype 3a, 14/284 – HCV4r genotype, 14/286 – HCV genotype 5a, 14/288 – genotype 6l.

The results obtained confirmed the ability of the “HEPA-BCD-test” reagent kit to detect genotypes A, B, C, D, E, F, G of the hepatitis B virus at a concentration of ~48 IU/ml and genotypes 1a, 1b, 2i, 3a, 4r, 5a, 6l hepatitis C virus at a concentration of ~100 IU/ml in 100 µl K2-EDTA plasma samples with an upper one-sided confidence interval of 95% exceeding the expected detection rate of 95%.

    - Analytical specificity: checking the effect of potentially interfering substances

The list of tested potentially interfering substances is given in section 8.3 of the Instructions.

Based on the results of the study, the following substances were classified as PCR inhibitors during the analysis:

- anticoagulants – heparin at a concentration of 0.15 IU/ml and sodium citrate at a concentration of 0.1 mM/ml. The use of heparin and sodium citrate as an anticoagulant when collecting peripheral blood is not allowed.

- heparin at a concentration of 1 IU/ml, used for anticoagulant therapy. The presence of heparin in the blood of patients on anticoagulant therapy can lead to unreliable PCR results, so it is recommended to draw blood from such patients before the next administration of the drug.

Other interfering substances in the indicated concentrations of interfering substances do not affect the results of qualitative detection and differentiation of hepatitis B virus, hepatitis C and D viruses using the “HEPA-BCD-test” test kit.

    - Analytical specificity: checking the effect of potential cross-reagents.

As part of the technical tests, the absence of false negative and false positive reactions was confirmed, respectively, with samples positive for HBV DNA and HCV and HDV RNA content and negative samples with the addition of DNA of the following pathogens:

- from the ATCC collection (American Type Culture Collection, USA): Propionibacterium acnes (ATCC® 29399™), Staphylococcus aureus subsp. aureus, strain Seattle 1945 (ATCC® 25923™), Staphylococcus epidermidis, FDA Strain PCI 1200 (ATCC® 12228™), Staphylococcus haemolyticus ATCC™ 29970™.

- from the collection "GKPM-OBOLENSK": Escherichia coli M-17 (strain number B-2929), Candida albicans NCTC 885-653 (strain number B-7940)

at a maximum concentration of 106 to 107 cells per ml. No nonspecific reactions were detected.

    - Precision in terms of repeatability and reproducibility:

      - The coefficient of variation under repeatability conditions is no more than 3%.

      - The coefficient of variation under kit reproducibility conditions is no more than 5%.

  - Clinical efficacy characteristics

    - Assessment of diagnostic sensitivity and diagnostic specificity

For clinical trials, 81 samples of human K2-EDTA blood plasma were used, obtained from a bank of residual aliquots formed during routine diagnostic and treatment practice from patients:

  - 41 samples of human K2-EDTA blood plasma, from patients with detected hepatitis C virus aged from 12 to 65 years (17 women and 24 men);

  - 30 samples of human K2-EDTA blood plasma, from patients with detected hepatitis B virus aged from 6 months to 85 years (from 14 women and 16 men)

  - 10 samples of human K2-EDTA blood plasma, from patients with detected hepatitis B and D virus aged from 18 to 65 years (from 4 women and 6 men),

having different stages of the corresponding disease.

This number of samples was collected in accordance with the recommendations of the International Guide CLSI EP09-A3, as well as in accordance with the requirements of GOST R 51352-2013.

To assess the interrater consistency of the study

clinical samples were tested with the study medical device in two series.

The following amplifiers were used to carry out the PCR study:

  - Detecting amplifier DTprime (NPO DNA-Technology LLC, Russia);

  - Amplifier CFX 96 (Bio-Rad, USA);

  - Rotor-Gene Q amplifier (Qiagen, Germany);

  - QuantStudio 5 amplifier (Thermo Fisher Scientific, USA). Of all tested samples of clinical material

10 samples are true positive for the NA content of hepatitis B and D viruses, 30 samples are true positive for the DNA content of the hepatitis B virus, and 41 samples are true positive for the RNA content of the hepatitis C virus. The results of qualitative detection of NA viruses for hepatitis B, C, D coincided with testing with the test kit of reagents "HEPA-BCD-test", as well as with the comparison kits "AmpliSens® HBV/HDV-FL" and

"AmpliSens® HCV-FL" for all tested samples of clinical material (no discordant samples were identified).

The results of studying the diagnostic characteristics of the studied set of reagents in relation to specific areas of the NA of hepatitis B, C, D viruses are shown in Table 4.

Table 4

| Defineanalyte | Number of observations with positive samples | Number of observations with negative samples | Diagnostic sensitivity with confidence95% probability | Diagnostic specificitywith confidence95% probability |
| --- | --- | --- | --- | --- |
| HBV DNA | 80 | 144 | 100% (95% CI: 95.49%- 100%) | 100% (95% CI: 97.47%- 100%) |
| HСV RNA | 82 | 142 | 100% (95% CI: 95.60%- 100%) | 100% (95% CI: 97.44%- 100%) |
| HDV RNA | 20 | 204 | 100% (95% CI: 83.16%- 100%) | 100% (95% CI: 98.21%- 100%) |

    - Cross-reactivity assessment

To assess cross-reactivity in clinical trials, the tested set of reagents “HEPA-BCD-test” also tested 35 samples that were truly negative, not containing the test analytes, but with confirmed, registered medical devices, presence of the following microorganisms: HIV-1 – 2 samples , adenovirus type 5 – 3 samples, varicella zoster virus – 1 sample, cytomegalovirus – 2 samples, Staphylococcus aureus – 5 samples, Epstein-Barr virus – 2 samples, hepatitis A virus – 2 samples, T-cell lymphotropic human virus type 2 – 2 samples, human herpes virus, type 6 – 2 samples, human papillomavirus – 3 samples, herpes simplex virus, type 1 – 4 samples, herpes simplex virus, type 2 – 3 samples.

No cross-reactivity was observed and no nonspecific reactions were detected.

    - Specificity assessment

The specificity of the “Kit of reagents for the simultaneous qualitative detection of NA viruses of hepatitis B, C, D using the multiplex RT-PCR “HEPA-BCD-test”, produced by TestGen LLC, was determined by testing negative content of NA viruses of hepatitis B, C, D samples from individual donors and pregnant women.

A total of 54 K2-EDTA blood plasma samples were tested using the “HEPA-BCD-test” reagent kit. 54 samples of K2-EDTA blood plasma turned out to be negative for the content of NA viruses of hepatitis B, C, D. When testing this panel of samples, the specificity of the HEPA-BCD test was 100.0 % (with a one-sided 95% confidence interval of 99.5%). The lower limit of the 95% confidence interval for specificity was determined by the method of Clopper-Pearson Confidience Interval; Clopper, C., & Pearson, E. (1934). The Use of Confidence or Fiducial Limits Illustrated in the Case of the Binomial. Biometrika, 26(4), 404-413. doi:10.2307/2331986).

To conduct a PCR study with the “HEPA-BCD-test” reagent kit, the amplifiers recommended by the manufacturer of the reagent kit under study were used:

- Detecting amplifier DTprime (NPO DNA-Technology LLC, Russia);

- Amplifier CFX 96 (Bio-Rad, USA);

- Rotor-Gene Q amplifier (Qiagen, Germany);

- QuantStudio 5 amplifier (Thermo Fisher Scientific, USA).

5. List of risks associated with the use of the reagent kit

The border risk zone includes the following hazards:

- Loss of functional properties of the reagents included in the kit due to transportation, storage or operation under inappropriate conditions;

- Contamination of clinical material with inhibitory substances in concentrations exceeding permissible limits;

- Contamination of reaction mixtures and samples of the test NA with contents from the PCS tube or amplification products;

- Conducting an analysis using a low quality NA sample (low concentration and/or poor purification);

- Failure to comply with the requirements for sample preparation, analysis and disposal due to work with the recruitment of unqualified personnel;

- Using a kit that is unsuitable for use (use after the expiration date or if the packaging is damaged).

No risks were identified in the area of the unacceptable zone.

The total residual risk of using the medical product “Reagent kit for simultaneous qualitative detection of NA viruses of hepatitis B, C, D by multiplex RT-PCR “HEPA-BCD-test”” is acceptable; the benefit of its use outweighs the risk.

6. Precautions when working with the kit

Class depending on the potential risk of use - 3, in accordance with the nomenclature classification of medical devices, approved by order of the Ministry of Health of the Russian Federation dated 06.06.2012 N 4n.

All components and reagents included in the “HEPA-BCD-test” reagent kit belong to hazard class 4 (low-hazard substances) in accordance with GOST 12.1.007-76 “SSBT. Harmful substances. Classification and general safety requirements."

The reagents included in the HEPA-BCD test kit have low vapor pressure and eliminate the possibility of inhalation poisoning.

The reagents included in the HEPA-BCD test kit are non-toxic because they are prepared by mixing individual non-toxic components.

Work with material that is contaminated or suspected of being contaminated is carried out in accordance with the requirements of SanPiN 3.3686-21 “Sanitary and epidemiological requirements for the prevention of infectious diseases”, MU “Organization of the work of laboratories using nucleic acid amplification methods when working with material containing microorganisms I– IV pathogenicity group" (MU 1.3.2569-09).

It is necessary to simultaneously provide and observe by personnel the rules of biological safety and the requirements for the organization and conduct of these works in order to prevent contamination of premises and equipment with nucleic acids and (or) amplicons of the studied samples.

The work should be carried out in a laboratory that performs molecular biological (PCR) studies of clinical material in compliance with sanitary and epidemiological rules SanPiN 2.1.3684-21 “Sanitary and epidemiological requirements for the maintenance of territories of urban and rural settlements, for water bodies, drinking water and drinking water supply , atmospheric air, soils, residential premises, operation of industrial and public premises, organization and implementation of sanitary and anti-epidemic (preventive)

events." Follow the recommendations set out in the MU

“Guidelines for disinfection, pre-sterilization cleaning and sterilization of medical devices” (MU 287-113), MU “Organization of the work of laboratories using nucleic acid amplification methods when working with material containing microorganisms of pathogenicity groups I–IV” (MU 1.3.2569 -09).

When working, you should always comply with the following requirements:

- remove unused reagents in accordance with SanPiN 2.1.3684-21 “Sanitary and epidemiological requirements for the maintenance of territories of urban and rural settlements, water bodies, drinking water and drinking water supply, atmospheric air, soil, residential premises, operation of industrial, public premises, organization and carrying out sanitary and anti-epidemic (preventive) measures";

- when removing waste after amplification (tubes containing PCR products), opening the tubes and splashing the contents is unacceptable, as this can lead to contamination of the laboratory area, equipment and reagents with amplification products;

- use the kit strictly for its intended purpose, according to these instructions;

- allow only specially trained personnel to work with the kit (a specialist with a higher medical education who has completed licensed specialization courses in working with pathogenic pathogens of groups I–II of pathogenicity and in PCR diagnostics, as well as a laboratory assistant with a secondary specialized medical education);

- do not use the kit after the expiration date;

- avoid contact with skin, NCSs and mucous membranes; In case of contact, immediately rinse the affected area with water and seek medical attention.

There are no necessary precautions regarding the influence of magnetic fields, external electrical influences, electrostatic discharges, pressure or pressure changes, overload, thermal ignition sources

provided.

The kit does not contain substances of human or animal origin that have a potential infectious nature, therefore precautions against any special, unusual risks when using or selling the product are not provided.

- Equipment and materials required when working with the reagent kit

Work with a set of reagents is carried out in work zone 3 (for preparing reactions) (MU 1.3.2569-09).

Equipment for multiplex RT-PCR:

- Biological safety box II and III protection class;

- Vortex;

- A set of electronic or automatic variable volume dispensers;

- Refrigerator from +2 °C to +8 °C with freezer not higher than -16 °C;

- Amplifier1 with real-time fluorescent detection through channels corresponding to FAM/Green, HEX/Yellow, ROX/Orange and Cy5/Red fluorophores, for example, CFX96 (BioRad, USA),

DTprime (DNA-Technology LLC, Russia), Rotor-Gene Q (Qiagen, Germany), QuantStudio 5 (Thermo Fisher Scientific, USA).

Materials and reagents not included in the product:

ATTENTION!When working with NA, it is necessary to use only disposable sterile plastic consumables that are specially marked “DNase-free” and “RNase-free”.

- Disposable tips with aerosol barrier up to 1000 µl, 200 µl, 20 µl and 10 µl (for example, Axygen, USA);

- Disposable sterile Eppendorf tubes of 1.5 or 2.0 ml;

- Thin-walled	disposable	test tubes	With	optically

1 Cyclers must be maintained, calibrated and used in accordance with the manufacturer's recommendations. Use of this kit in an uncalibrated instrument may affect the performance of the reagent kit.

transparent lid (when using plate-type amplifiers) or optically transparent walls (when using rotary-type amplifiers) for PCR: PCR tubes with a volume of 0.1 or 0.2 ml2, or tubes for PCR with a volume of 0.1 or 0.2 ml in strips, or PCR plates with optically transparent film (for example, Axygen, USA), compatible with the amplifier used;

- Gown and disposable gloves without talc;

- Container with disinfectant solution;

- Workplace racks for tubes with a volume of 0.1 or 0.2 ml or for strip tubes with a volume of 0.1 or 0.2 ml (for example, InterLabService LLC, Russia);

- Kit for isolating NA from blood plasma K2-EDTA (see paragraph.

8.2 Instructions).

- Samples analyzed Type of sample analyzed

The material for RT-PCR is nucleic acid samples isolated from blood plasmaK2-EDTA.

8.1. Procedure for obtaining clinical material

ATTENTION! Before starting work, you should familiarize yourself with the methodological recommendations “Taking, transporting, storing clinical material for PCR diagnostics”, developed by the Federal Budgetary Institution Central Research Institute of Experimental Inspection of Rospotrebnadzor, Moscow, 2012.

The collection of clinical material and its packaging is carried out by an employee of a medical organization trained in the requirements and rules of biological safety when working and collecting material suspected of being infected with microorganisms of pathogenicity group II.

Collection of material for research

Peripheral blood is collected in the morning on an empty stomach into a test tube (vacuum tube) containing a 4 or 6 ml EDTA-K2 solution as an anticoagulant. Immediately after drawing blood, invert the tube 3-4 times to mix the blood with the EDTA-K2 solution.

ATTENTION! The use of heparin and

2 Make sure the PCR tubes are compatible with the cycler you are using.

sodium citrate as an anticoagulant.

ATTENTION! The presence of heparin in the blood of patients on anticoagulant therapy can lead to unreliable PCR results, so it is recommended to draw blood from such patients before the next administration of the drug.

Conditions for transportation and storage of the initial clinical material - blood:

- at temperatures from +2 °C to +8 °C – no more than 6 hours;

- at room temperature – no more than 2 hours. Do not freeze blood.

Within 2 hours (when stored at room temperature) or 6 hours (when stored at temperatures from +2 ° C to +8 ° C) after collecting the material, plasma should be collected, for which the tube with blood is centrifuged at 800–1600 g for 20 minutes at room temperature. After centrifugation, transfer the upper fraction (plasma) into separate plastic tubes with a volume of 1.5 or 2.0 ml, free from DNase and RNase.

Conditions for transportation and storage of K2-EDTA blood plasma:

Plasma can be stored at temperatures from +2 °C to

+8 °C for up to 5 days, at temperatures from -18 °C to -22 °C - up to 3 months, at temperatures -70 °C - for a long time.

ATTENTION!Avoid repeated freezing and thawing of plasma samples.

To isolate NA, use at least 100 μl of plasma. Increasing the analytical sensitivity of the kit is possible by using a larger volume of plasma, if provided by the kit used for NA isolation, as well as reducing the elution volume (optimally 25 µl).

In exceptional cases, the use of blood serum is allowed, and the reduction in the clinical sensitivity of the reagent kit while maintaining analytical sensitivity should be taken into account. Store the serum at a temperature from +2 °C to +8 °C for up to 5 days, at a temperature from -18 °C to -22 °C for up to 3 months, at a temperature of -70 °C for a long time.

Material pre-treatment

No preparation required.

Recording, storage, transfer and transportation of clinical material suspected of the presence of hepatitis virus must be carried out in accordance with the current sanitary and epidemiological rules for the safety of working with microorganisms of pathogenicity (hazard) groups I–II (SP 1.3.3118-13), the current sanitary rules on the procedure for recording, storage, transfer and transportation of microorganisms of pathogenicity groups I–IV.

Disposal of clinical material (class B), as extremely epidemiologically hazardous waste, is carried out in accordance with SanPiN 2.1.3684-21.

8.2 Procedure for obtaining a human NA sample isolated from K2-EDTA blood plasma

To isolate a human NA sample from K2-EDTA blood plasma, it is recommended to use the following reagent kits:

- A set of reagents for the isolation of DNA/RNA from clinical material “NA-Extra” according to TS 21.20.23-013-97638376-2019 produced by TestGen LLC, Russia (registration certificate: RZN 2021/15428 dated 09.24.2021).

During the NA isolation procedure, it is necessary to strictly follow the protocol and the requirements of the instructions for the reagent kit used.

To each test K2-EDTA plasma sample, 10 µl of ICS from the reagent kit should be added before isolation

"HEPA-BCD test".

The NCS sample also undergoes the procedure of NC isolation in a volume of 100 μl with the addition of 10 μl of ICS. If the manufacturer's instructions for the NA isolation reagent kit require the use of a larger volume of sample, the NCS volume should be adjusted to the required volume with saline or TE buffer.

Conditions for possible storage of analyzed samples

NA:

- at +2 ... +8 °C no more than 4 hours (recommended),

- at -18... -22 °C for no more than a week,

- at a temperature not exceeding -80 °C for no more than a year.

8.3. Interfering substances and restrictions on the use of the analyzed material

The effect of potentially interfering substances on the performance of the HEPA-BCD test kit has been tested against potentially interfering substances that may be encountered during normal use of the reagent kit.

"HEPA-BCD test", and presumably influence the abilityset of reagentsproduce reliable results.

Interfering substances may come from the following external and internal sources:

- substances used in the treatment of the patient (for example, drugs);

- substances found in specific types of samples - in this case, contamination of a clinical sample with blood hemoglobin can inhibit PCR if insufficient purification during the NA isolation procedure;

- substances encountered during the collection of clinical material –in this case,anticoagulants.

The studied concentrations of interfering substances are given in Table 5.

Table 5

| Interfering substances | Maximumconcentration |
| --- | --- |
| Endogenous interfering substances |  |
| Hemoglobin | 260 µg/ml |
| Heparin (anticoagulant) | 0.15 IU/ml |
| Sodium citrate (anticoagulant) | 0.1 mm/ml |
| EDTA-K2 (anticoagulant) | 0.5 mm/ml |
| Cholesterol | 150 mg/dl |
| Triglycerides | 250 mg/dl |
| Exogenous interfering substances |  |
| During anticoagulant therapy |  |
| Heparin | 1 IU/ml |
| Drugs prescribed for viral hepatitis B, C, D |  |
| Interferon alpha | 1000 IU/ml |
| Pegylated interferon alfa | 0.036 µg/ml |
| Drugs prescribed for viral hepatitis B |  |

| Lamivudine | 0.02 mg/ml |
| --- | --- |
| Entecavir | 0.1*10-3 mg/ml |
| Telbivudin | 0.12 mg/ml |
| Drugs prescribed for viral hepatitis C |  |
| Ribavirin | 0.04 mg/ml |
| Narlaprevir | 0.02 mg/ml |
| Paritaprevir | 0.015 mg/ml |
| Dasabuvir | 0.05 mg/ml |
| Sofosbuvir | 0.08 mg/ml |
| Daclatasvir | 0.012 mg/ml |
| Ledipasvir | 0.018 mg/ml |
| Ombitasvir | 0.0025 mg/ml |
| Drugs prescribed for viral hepatitis D |  |
| Bulevirtide | 0.0004 mg/ml |

Based on the results of the study, the following substances were classified as PCR inhibitors during the analysis:

- anticoagulants – heparin at a concentration of 0.15 IU/ml and sodium citrate at a concentration of 0.1 mM/ml. The use of heparin and sodium citrate as an anticoagulant when collecting peripheral blood is not allowed.

- heparin at a concentration of 1 IU/ml, used for anticoagulant therapy. The presence of heparin in the blood of patients on anticoagulant therapy can lead to unreliable PCR results, so it is recommended to draw blood from such patients before the next administration of the drug.

Restrictions on the use of the analyzed material:

- the analyzed material cannot be used if storage and transportation conditions are violated (temperature, duration, repeated freezing-thawing);

- Samples contaminated with foreign biological material must not be used.

the presence of heparin in the blood of patients on anticoagulant therapy can lead to unreliable PCR results, therefore, blood sampling from such patients is recommended before the next administration drug.

- Preparing test kit components

Installation, installation, configuration, calibration of the medical device for commissioning is not required.

ATTENTION!When working with NA, it is necessary to use only disposable sterile plastic consumables that are specially marked “DNase-free” and “RNase-free”. Be sure to use a separate aerosol barrier tip for each reaction component.

ATTENTION!The components of the reaction mixture should be mixed immediately before analysis.

Before preparing reaction mixtures, it is necessary to wet clean the PCR box, as well as the equipment and materials located in it, using disinfectants suitable for use in PCR laboratories, and turn on the UV lamp for 20–30 minutes.

- Thoroughly mix the contents of the tubes with the NA isolated for analysis, PCR buffer, a mixture of oligonucleotides, NCS and PCS, inverting each tube 10 times or vortexing at low speed for 3-5 seconds, then sediment the drops from the caps of the tubes by short centrifugation.

- Select the required number of tubes (with optically transparent lids or walls - depending on the type of detection amplifier used) with a volume of 0.1 or 0.2 ml for PCR based on: the number of samples being studied3

+ 1 PCS + 1 NCS, or take a PCR tablet.

10. Conducting analysis

PCR research consists of the following stages:

- RT-PCR preparation;

- Reverse transcription of RNA and PCR amplification of DNA with hybridization-fluorescence detection in real time;

- Interpretation of results.

3To improve accuracy, it is recommended that each sample be analyzed in duplicate.

A) Preparation of RT-PCR

(produced in the pre-PCR ZONE room for unpacking reagents and preparing for PCR amplification)

The total reaction volume is 25 µl.

ATTENTION!It is forbidden to change the reaction volume.

For	preparations	reactionary	mixtures	on	1	reactionnecessary:

  - PCR buffer – 5 µl,

  - Mixture of oligonucleotides – 5 µl,

  - Sample (test sample NA, PCS, NCS) – 15 µl.

Reaction tubes must be prepared in the following order:

- Label 0.1 or 0.2 ml tubes or PCR plate.

- In a separate disposable sterile tube type

“Eppendorf” with a volume of 1.5 or 2.0 ml, prepare a reaction mixture: (n+3)x5 μl of PCR buffer and (n+3)x5 μl of a mixture of oligonucleotides, where n is the number of samples being tested.

- Add 10 µl of the prepared reaction mixture to each PCR tube.

- Add 15 μl of isolated NA into the appropriate test tubes for the samples being studied. The NA preparation is not added to the test tubes for PCS and NCS.

- Add 15 µl of PCS into the appropriate tube.

- Add 15 μl of NCS, which has passed the stage of NC isolation, into the appropriate test tube.

- To remove droplets from the walls, centrifuge the tubes for 1–3 seconds in a microcentrifuge vortex.

B) Reverse transcription of RNA and PCR amplification of DNA with hybridization-fluorescence detection in real time;

(produced in the PCR ZONE - the room for PCR amplification)

- Place the tubes in the reaction module of the real-time PCR device. It is recommended to install the tubes in the center of the thermoblock to ensure uniform pressure on the tubes with the heating lid.

- Program the device to perform

appropriate RT-PCR program and fluorescent signal detection, following the instructions for the device used. The RT-PCR protocol is listed in Table 6.

Table 6 – RT-PCR protocol

| Stage | Temperature, °C | Time, min.:sec. | Detection channels | Total cycles |
| --- | --- | --- | --- | --- |
| 1 | 52 | 40:00 | - | - |
| 2 | 95 | 02:00 | - | - |
| 3 | 95 | 00:05 | - | 5 |
|  | 60 | 00:15 | - |  |
|  | 67 | 00:30 | - |  |
| 4 | 95 | 00:05 | - | 45 |
|  | 60 | 00:15 | FAM/Green, HEX/Yellow, ROX/Orange,Cy5/Red |  |
|  | 67 | 00:30 | - |  |

ATTENTION!When using the QuantStudio 5 amplifier, it is necessary to adjust the optical filters before starting the amplification protocol (since without this setting, the VIC signal can be recorded by the ROX channel in the case of ΔRn > 100,000, which can lead to false positive results)! To do this, in the “Method” tab, click the “Action” button, and then select from the pop-up menu

“Optical filter settings”, where in the “PCR Filter” section leave only the following filter combinations: x1 – m1, x2 – m2, x4 – m4, x5 – m5, x6 – m6.

- Indicate the number and identifiers of samples, mark the location of the tubes on the thermoblock matrix in accordance with their installation.

- Make sure that the FAM/Green, HEX/Yellow, ROX/Orange and Cy5/Red detection channels are enabled in the optical measurement parameters of the amplification program.

- Run RT-PCR with fluorescent signal detection.

- After completing the program, begin analyzing the results.

  - Registration and interpretation of results

Registration of results is carried out upon completion of RT-PCR automatically using the software of the device used.

Recommendations for setting the threshold line

For amplifiers of any model, the threshold line is set individually for each detection channel at a level corresponding to 5–10% of the maximum fluorescence level obtained for the positive control sample in the last amplification cycle.

The interpretation of the results is carried out based on the Ct values for the detection channels corresponding to the fluorophores FAM/Green, HEX/Yellow, ROX/Orange and Cy5/Red (Table 1). Only Ct values obtained at the PCR stage with fluorescent detection are taken into account (that is, corresponding to stage 4 - see Table 6).

First, it is necessary to evaluate the progress of the reaction and the Ct values in the control samples. Interpretation of the results in the studied samples begins only after the correct passage of the PCS and NCS.

ATTENTION!In the case of using the Rotor-Gene 6000, Rotor-Gene 3000, Rotor-Gene Q and similar amplifiers, activate the “Dynamic” functions. background" (Dynamic Tube), "Correct. slope" (Noise slope correction), set the value to 10% in the section

Outlier Removal.

Interpretation of results in control samples

For NCS and PCS, the following results should be obtained (Table 7).

Table 7 – Study results for NCS and PCS

| Controlsample | Ct values for detection channels corresponding to fluorophores |  |  |  |
| --- | --- | --- | --- | --- |
|  | FAM/Green | HEX/Yellow | ROX/Orange | Cy5/Red |
| NCS | > 35 or absent | ≤ 32 | > 35 or absent | > 35 or absent |
| PCS | ≤30 | ≤30 | ≤30 | ≤30 |

If values for the NCS are obtained that differ from those indicated in Table 7, the results of the entire production series are considered unreliable. In this case, it is necessary to carry out

special measures to eliminate possible contamination.

If values for the PCS are obtained that differ from those indicated in Table 7, amplification of the entire batch of samples must be repeated. If PCS values are repeatedly obtained that differ from those indicated in Table 7, it is necessary to replace the reagents.

Interpretation of results in clinical samples studied

The principles for interpreting the results are reflected in Table 8.

If for a sample an increase in the fluorescence of a specific product is recorded through the FAM/Green, ROX/Orange or Cy5/Red channels before cycle 15 (Ct < 15), then this indicates a high initial concentration of NA of the corresponding hepatitis virus. In this case, it is possible to obtain a false negative result for the hepatitis virus, the NA of which is present in low concentrations. To exclude false negative results, it is recommended to repeat RT-PCR/PCR of the isolated NA preparation or use reagent kits for separate detection of NA of the corresponding hepatitis.

The reason for obtaining an invalid result may be a low concentration of NA, the presence of inhibitors in the NA preparation obtained from clinical material, incorrect implementation of the analysis protocol, non-compliance with the RT-PCR temperature regime, etc.

The reason for obtaining a questionable result may be a low concentration of the virus in the clinical sample.

In the case of an invalid and questionable result, a conclusion is not issued; it is necessary to re-take biomaterial from the patient and re-run the analysis. In this case, for doubtful results, it is recommended to isolate NC from a larger volume of plasma.

Table 8 – Principle of interpretation of results in the studied clinical samples

| Ct values ​​by detection channels, corresponding fluorophores (target analytes are indicated in parentheses) |  |  |  | Result |
| --- | --- | --- | --- | --- |
| FAM/ Green (HBV) | ROX/ Orange(HCV) | Cy5/ Red (HDV) | HEX/ Yellow (ICS) |  |
| - | - | - | ≤ 32 | NA of hepatitis B, C and D viruses was not detected |
| - | - | - | > 32 | the result is invalid |
| ≤ 35 |  |  | not counted | Virus NA corresponding to the channel was detected |
| > 35 |  |  | not counted | test result is questionable for the target corresponding to the channel |

Designations: “not accounted for.” – the result is not taken into account during interpretation; “-” – there is no fluorescence signal.

If a questionable result is repeated, repeat the test using a set of reagents from another manufacturer or another method.

  - Conditions for storage, transportation and operation of the reagent kit

Storage

The “HEPA-BCD-test” reagent kit in the manufacturer’s packaging should be stored at a temperature from minus 18 to minus 22 °C throughout the entire shelf life of the kit; storage at a temperature from 2 to 8 °C is allowed for no more than 30 days.

Freezing/thawing of the HEPA-BCD test kit is allowed no more than 10 times.

After opening, store under the same conditions as the reagents before opening.

A set of reagents stored in violation of the regulated regime cannot be used.

Transportation

The “HEPA-BCD-test” reagent kit should be transported by all types of transport in covered vehicles in accordance with the transportation rules in force for this type of transport.

Transport at temperatures from minus 18 to minus 22 °C during the entire shelf life of the kit. Transportation is allowed at temperatures from 2 to 8 °C for up to 30 days, or at temperatures from 15 to 25 °C for no more than 5 days.

Atmospheric pressure is not controlled, as it does not affect the quality of the product.

To ensure compliance with transportation conditions throughout the entire transportation period, the set of reagents is placed in a reusable polyurethane foam thermal container for temporary storage and transportation with prepared cold elements. The type, volume and quantity of cold elements placed in a thermal container with transported reagent kits, as well as the volume of the thermal container, are selected depending on the duration and conditions of transportation.

Reagent kits transported in violation of the temperature regime cannot be used.

Best before date

The shelf life of the “HEPA-BCD-test” reagent kit is 12 months from the date of acceptance by the manufacturer’s quality control department, subject to all conditions of transportation, storage and operation. Expired reagent kits cannot be used.

Shelf life of opened kit components

12 months from the date of acceptance by the manufacturer’s quality control department, subject to storage at a temperature from minus 18 to minus 22 ° C.

Shelf life of kit components prepared for use

One hour, subject to conditions that prevent the components from drying out, as well as contamination with foreign biological material.

  - Disposal

Reagent kits that have become unusable, including due to the expiration date, must be disposed of in accordance with the requirements of SanPiN 2.1.3684-21 “Sanitary and epidemiological requirements for the maintenance of territories of urban and rural settlements, for water bodies, drinking water and drinking water”. water supply, atmospheric air, soil, residential premises, operation of industrial and public premises, organization and implementation of sanitary and anti-epidemic (preventive) measures.”

In accordance with the classification of medical waste, the kits belong to class A (epidemiologically safe waste, similar in composition to solid household waste). Unused reagents in accordance with SanPiN 2.1.3684-21

“Sanitary and epidemiological requirements for the maintenance of territories of urban and rural settlements, for water bodies, drinking water and drinking water supply, atmospheric air, soils, residential premises, operation of industrial and public premises, organization and implementation of sanitary and anti-epidemic (preventive) measures” are collected in disposable labeled packaging of any color (except yellow and red).

The test tubes and materials remaining after completion of the work are disposed of in accordance with MU “Guidelines for disinfection, pre-sterilization cleaning and sterilization of medical devices” (MU 287-113).

Liquid components (reagents, reagents) are destroyed by draining them into the sewer with preliminary dilution of the reagent with tap water 1:100 and removal of the remaining packaging as industrial or household waste.

Consumer packaging of the “HEPA-BCD-test” reagent kit is subject to mechanical destruction with removal of the remainder as industrial or household waste.

Personnel carrying out the destruction of a set of reagents must comply with the safety rules for carrying out a particular method of destruction.

  - Warranty obligations, contacts

The manufacturer guarantees the quality and safety of the reagent kit "HEPA-BCD-test" during the shelf life in case of compliance with the requirements of transportation and storage of products, as well as in case of compliance with the rules of operation.

In case of complaints about reagent kit quality, undesirable events or incidents, submit information to:

Limited Liability Company “TestGene”

(TestGene, LLC),

9 44th Inzhenerny Proezd, office 13, Ulyanovsk 432072

Tel.: +7 (499) 705-03-75

www.testgen.ru

Technical Support Service:

Tel.: +7 927 981 58 81

E-mail: help@testgen.ru